Appendix B : Guidance on the construction of an anechoic shielded chamber

B.1	Construction of the anechoic shielded chamber

Free field measurements can be simulated in an anechoic shielded chamber where the walls are coated with RF absorbers. Fig. App.B-1 shows the requirements for shielding loss and wall return loss of such a chamber. As dimensions and characteristics of usual absorber materials are critical below 100 MHz (height of absorbers approximately 1m, reflection attenuation < 20 dB) such a room is preferably suitable for measurements above 100 MHz.

The anechoic shielded chamber shall have dimensions of at least 5x5x10 metres. Ceilings and walls are covered with pyramidically shaped absorbers of about 1 meter. The floor is covered with floor absorbers, on which it is possible to walk. As the internal dimensions are 3 m x 3 m x 8 m, a measuring distance of max. 5 m length is available in the middle of the chamber. Fig. App.B-2 shows a possible layout of the chamber.

The wall absorbers and floor absorbers reject reflections so that the antenna height need not be changed.

B.2	Calibration of the anechoic shielded chamber

The calibration of the chamber should be checked, measuring the field strength as a function of the distance from the radiation antenna, for various frequencies over the measuring range. The calibration curve is used to correct the readings obtained during the measurements.

An example of a calibration curve is given in Fig. App.B-3. In the range of 30 MHz to 10 GHz, the disturbing influences due to imperfections of the anechoic shielded chamber, may be in the order of 1 to 2 dB at a measuring distance of 3 metre.

.G.1110figs\F-XX-B1.EPS;16.499 cm;23.547 cm;EPS

.G.1110figs\F-XX-B2.EPS;16.499 cm;23.547 cm;EPS

.G.1110figs\F-XX-B3.EPS;16.499 cm;24.698 cm;EPS